## SummaryForFigure
Acquisition Efficiency of high-resolution UHF BOLD fMRI Sequences, computed as resolved voxels per unit time (matrix size/second)
Sequence | Publication | resolution [mm] |  |  | FOV [mm] |  |  | matrix size |  |  | inter-leaves | TR / interleaf | volume TR | acquisition efficiency | comment
|  | x | y | z | x | y | z | x | y | z |  | [s] | [s] | [1000 voxels/s]
3D GRASE (Park 2021) | (Park et al., 2021) | 0.8 | 0.8 | 0.8 | 89.6 | 22.4 | 28.8 | 112 | 28 | 36 | 1 | 1 | 3 | 37.632 | with nonlinear k-t recon, half of that w/o prior info
3D T2-prep FLASH (Hua 2014) | (Hua et al., 2014) | 2.5 | 2.5 | 2.5 | 210 | 210 | 137.5 | 84 | 84 | 55 | 532.2222 | 0.0036 | 2.3 | 168.7304 | GRAPPA 3x3, whole 3D volume read out in single shot
2D SE EPI (Kemper 2015) | (Kemper et al., 2015) | 0.8 | 0.8 | 0.8 | 120 | 120 | 16 | 150 | 150 | 16 | 1 | 2 | 2 | 180
bSSFP (Ehses 2018) | (Ehses and Scheffler, 2018) | 0.9 | 0.9 | 0.9 | 189 | 189 | 14.4 | 210 | 210 | 16 | 314 | 0.01 | 3.14 | 224.7134
2D GRE EPI (Nasr 2016) | (Nasr et al., 2016) | 1 | 1 | 1 | 192 | 192 | 44 | 192 | 192 | 44 | 1 | 3 | 3 | 540.672 | GRAPPA 4, body gradient (70/200), otherwise sequence as in Polimeni et al., 2010
2D GRE EPI (theory) |  | 0.95 | 0.95 | 0.95 | 220 | 220 | 38 | 231.5789 | 231.5789 | 40 | 1 | 0.07 | 2.8 | 766.1258 | 40/200 gradient system, using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total
2D Spiral (Kasper 2022) | (Kasper et al., 2022) | 0.8 | 0.8 | 0.9 | 230 | 230 | 32.4 | 287.5 | 287.5 | 36 | 1 | 3.3 | 3.3 | 901.7045 | R=4, but resolution defined by k_max (not area, therefore sqrt(4/pi) factor difference)
3D GRE EPI (Poser 2010, R=9) | (Poser et al., 2010) | 1 | 1 | 1 | 180 | 180 | 104 | 180 | 180 | 104 | 64 | 0.05 | 3.2 | 1053 | 3x3 undersampling, 25% slice oversampling
SMS GRE EPI (HCP 7T, R=10) | HCP Resting State 7T (T. Vu et al., 2017) | 1.6 | 1.6 | 1.6 | 208 | 208 | 136 | 130 | 130 | 85 | 1 | 1 | 1 | 1436.5 | SMS 5, GRAPPA 2, Fourier 7/8 see https://www.humanconnectome.org/hcp-protocols-ya-7t-imaging
3D GRE EPI (theory) |  | 0.95 | 0.95 | 0.95 | 220 | 220 | 76 | 231.5789 | 231.5789 | 80 | 40 | 0.07 | 2.8 | 1532.2517 | like 2D, but Rz=2 (total R=8); 40/200 gradient system, using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total
3D GRE t-Hex Spiral (Engel 2022) | (Engel et al., 2022) | 0.6 | 0.6 | 2 | 240 | 240 | 120 | 400 | 400 | 60 | 86 | 0.056 | 4.816 | 1993.3555 | short TE 15 ms, effective R=8, proposed, no fMRI data yet
SMS GRE EPI (Vu 2017, HCP Pilot, R=9) | HCP pilot (T. Vu et al., 2017) | 0.9 | 0.9 | 0.9 | 210 | 210 | 135 | 233.3333 | 233.3333 | 150 | 1 | 3.733 | 3.733 | 2187.6953 | SMS 3, GRAPPA 3, Fourier 5/8 see https://www.humanconnectome.org/hcp-protocols-ya-7t-imaging
3D GRE EPI (van der Zwaag 2018, R=12) | (van der Zwaag et al., 2018) | 0.9 | 0.9 | 2 | 200 | 200 | 120 | 222.2222 | 222.2222 | 60 | 20 | 0.055 | 1.1 | 2693.6027 | 4x3 undersampling, 25% slice oversampling
3D GRE t-Hex Spiral-in (Engel 2021) | (Engel et al., 2021) | 1.43 | 1.43 | 1.43 | 240 | 240 | 120 | 167.8322 | 167.8322 | 83.9161 | 13 | 0.0662 | 0.8606 | 2746.5928 | short TE 15 ms, proposed, no fMRI data yet
3D GRE EPI (theory, head gradient) | Head Gradient Specs Feinberg 2021 | 0.525 | 0.525 | 0.525 | 220 | 220 | 42 | 419.0476 | 419.0476 | 80 | 40 | 0.07 | 2.8 | 5017.1688 | 200/900 head-gradient system (Feinbergatron), using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total

## AllReferences
Acquisition Efficiency of high-res UHF BOLD fMRI Sequences, computed as resolved voxels per unit time (matrix size/second)
Sequence | Publication | resolution [mm] |  |  | FOV [mm] |  |  | matrix size |  |  | inter-leaves | TR / interleaf | volume TR | acquisition efficiency | comment
|  | x | y | z | x | y | z | x | y | z |  | [s] | [s] | [1000 voxels/s]
3D GRASE | (Park et al., 2021) | 0.8 | 0.8 | 0.8 | 89.6 | 22.4 | 28.8 | 112 | 28 | 36 | 1 | 1 | 3 | 37.632 | with nonlinear k-t recon, half of that w/o prior info
3D T2-prep FLASH (multi-echo) | (Pfaffenrot et al., 2021) | 0.75 | 0.75 | 0.75 | 189 | 189 | 24 | 252 | 252 | 32 | 96 | 0.50125 | 48.12 | 42.2304 | multi-echo (4 echoes), Ry=3, 3 segments in y, 32 segments in z (Rz=1, because of slab FOV), TR_GRE 15.6 ms
2D SE EPI (double-echo) | (Han et al., 2021) | 0.8 | 0.8 | 0.8 | 120 | 120 | 12.8 | 150 | 150 | 16 | 3 | 2 | 6 | 60 | sliding window recon allows 1 image per shot (x3 efficiency, but temporal correlation)
EPTI | (Wang et al., 2021) | 0.9 | 0.9 | 0.9 | 266.6667 | 160 | 31.1111 | 240 | 144 | 28 | 3 | 3 | 9 | 107.52 | multiple echoes reconstructed per TR
3D T2-prep FLASH (no PF) | (Hua et al., 2014) | 2.5 | 2.5 | 2.5 | 210 | 210 | 137.5 | 84 | 84 | 55 | 532.2222 | 0.0036 | 2.3 | 168.7304 | GRAPPA 3x3, whole 3D volume read out in single shot
2D SE EPI | (Kemper et al., 2015) | 0.8 | 0.8 | 0.8 | 120 | 120 | 16 | 150 | 150 | 16 | 1 | 2 | 2 | 180
3D GRE EPI (Hendriks 2020, fMRI) | (Hendriks et al., 2020) | 0.8 | 0.8 | 0.8 | 51 | 163 | 28 | 63.75 | 203 | 35 | 44.2593 | 0.054 | 2.39 | 189.5162 | CAIPI 7x1 y4, used for fMRI time series
3D GRE EPI (Hendriks 2020, sub-second) | (Hendriks et al., 2020) | 0.95 | 0.95 | 0.95 | 65 | 176 | 15 | 68.4211 | 184 | 15.7895 | 16.4815 | 0.054 | 0.89 | 223.3496 | CAIPI 8x1 y3, proposed for sub-second fMRI, no time series in article
bSSFP | (Scheffler and Ehses, 2015) | 1 | 1 | 1 | 172 | 132 | 20 | 172 | 132 | 20 | 353.2164 | 0.0038 | 3.02 | 150.3576 | R = 3, 6/8 Partial Fourier, at 9.4T
bSSFP | (Ehses and Scheffler, 2018) | 0.9 | 0.9 | 0.9 | 189 | 189 | 14.4 | 210 | 210 | 16 | 314 | 0.01 | 3.14 | 224.7134
2D GRE EPI (R=3) | (Heidemann et al., 2012) | 0.75 | 0.75 | 0.75 | 162 | 166.5 | 23.25 | 216 | 222 | 31 | 1 | 3.3 | 3.3 | 450.4582 | GRAPPA 3, enables reduced FOV imaging
2D GRE EPI (R=3) | (Polimeni et al., 2010) | 1 | 1 | 1 | 192 | 192 | 34 | 192 | 192 | 34 | 1 | 2.5 | 2.5 | 501.3504 | GRAPPA 3, head-insert gradient (80/400)
2D GRE EPI | (Nasr et al., 2016) | 1 | 1 | 1 | 192 | 192 | 44 | 192 | 192 | 44 | 1 | 3 | 3 | 540.672 | GRAPPA 4, body gradient (70/200), otherwise sequence as in Polimeni et al., 2010
multiline bSSFP | (Ehses and Scheffler, 2018) | 0.9 | 0.9 | 0.9 | 189 | 189 | 14.4 | 210 | 210 | 16 | 121 | 0.01 | 1.21 | 583.1405
3D GRE EPI (Hendriks 2020, Anat) | (Hendriks et al., 2020) | 0.5 | 0.5 | 0.5 | 240 | 186 | 150 | 480 | 364 | 300 | 1058.3333 | 0.072 | 76.2 | 687.874 | CAIPI 7x1 y3, used for anatomical T2*
2D GRE EPI (theoretical) |  | 0.95 | 0.95 | 0.95 | 220 | 220 | 38 | 231.5789 | 231.5789 | 40 | 1 | 0.07 | 2.8 | 766.1258 | 40/200 gradient system, using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total
3D T2-prep FLASH (PF) | (Hua et al., 2014) | 1.5 | 1.5 | 1.6 | 210 | 210 | 134.4 | 140 | 140 | 84 | 540 | 0.0031 | 1.86 | 885.1613 | PF 5/8 (in y and z), GRAPPA 3x3, whole 3D volume read out in single shot
2D Spiral | (Kasper et al., 2022) | 0.8 | 0.8 | 0.9 | 230 | 230 | 32.4 | 287.5 | 287.5 | 36 | 1 | 3.3 | 3.3 | 901.7045 | R=4, but resolution defined by k_max (not area, therefore sqrt(4/pi) factor difference)
2D GRE Spiral (theoretical) |  | 0.85 | 0.85 | 0.85 | 220 | 220 | 34 | 258.8235 | 258.8235 | 40 | 1 | 0.07 | 2.8 | 956.9946 | would also work for EPI w/ 80/200 gradient system, using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total
3D GRE EPI (Poser 2010, R=9) | (Poser et al., 2010) | 1 | 1 | 1 | 180 | 180 | 104 | 180 | 180 | 104 | 64 | 0.05 | 3.2 | 1053 | 3x3 undersampling, 25% slice oversampling
2D GRE EPI (theoretical, long readout) |  | 0.55 | 0.55 | 0.55 | 220 | 220 | 16.5 | 400 | 400 | 30 | 1 | 0.12 | 3.6 | 1333.3333 | 80/200 gradient system, using TE=min (55ms) for TAQ, 5ms for excitation, TAQ 100 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 120 ms TR total
SMS GRE EPI (HCP 7T, R=10) | HCP Resting State 7T (T. Vu et al., 2017) | 1.6 | 1.6 | 1.6 | 208 | 208 | 136 | 130 | 130 | 85 | 1 | 1 | 1 | 1436.5 | SMS 5, GRAPPA 2, Fourier 7/8 see https://www.humanconnectome.org/hcp-protocols-ya-7t-imaging
3D GRE EPI (theoretical, R=8) |  | 0.95 | 0.95 | 0.95 | 220 | 220 | 76 | 231.5789 | 231.5789 | 80 | 40 | 0.07 | 2.8 | 1532.2517 | like 2D, but Rz=2 (total R=8); 40/200 gradient system, using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total
3D GRE t-Hex EPI (0.7 mm) | (Engel et al., 2022) | 0.7 | 0.7 | 2 | 240 | 240 | 120 | 342.8571 | 342.8571 | 60 | 50 | 0.086 | 4.3 | 1640.2468 | effective R = 8 proposed, no fMRI data yet
3D GRE t-Hex Spiral-in (2 mm) | (Engel et al., 2021) | 2.05 | 2.05 | 2.05 | 240 | 240 | 120 | 117.0732 | 117.0732 | 58.5366 | 9 | 0.045 | 0.405 | 1981.0121 | proposed, no fMRI data yet
3D GRE t-Hex Spiral | (Engel et al., 2022) | 0.6 | 0.6 | 2 | 240 | 240 | 120 | 400 | 400 | 60 | 86 | 0.056 | 4.816 | 1993.3555 | short TE 15 ms, effective R=8, proposed, no fMRI data yet
SMS GRE EPI (HCP Pilot, R=9) | HCP pilot (T. Vu et al., 2017) | 0.9 | 0.9 | 0.9 | 210 | 210 | 135 | 233.3333 | 233.3333 | 150 | 1 | 3.733 | 3.733 | 2187.6953 | SMS 3, GRAPPA 3, Fourier 5/8 see https://www.humanconnectome.org/hcp-protocols-ya-7t-imaging
3D GRE EPI (theoretical, long readout) |  | 0.55 | 0.55 | 0.55 | 220 | 220 | 33 | 400 | 400 | 60 | 30 | 0.12 | 3.6 | 2666.6667 | as 2D, but Rz=2 (total R=8); 80/200 gradient system, using TE=min (55ms) for TAQ, 5ms for excitation, TAQ 100 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 120 ms TR total
3D GRE EPI (van der Zwaag 2018, R=12) | (van der Zwaag et al., 2018) | 0.9 | 0.9 | 2 | 200 | 200 | 120 | 222.2222 | 222.2222 | 60 | 20 | 0.055 | 1.1 | 2693.6027 | 4x3 undersampling, 25% slice oversampling
3D GRE t-Hex Spiral-in (1.4 mm) | (Engel et al., 2021) | 1.43 | 1.43 | 1.43 | 240 | 240 | 120 | 167.8322 | 167.8322 | 83.9161 | 13 | 0.0662 | 0.8606 | 2746.5928 | short TE 15 ms, proposed, no fMRI data yet
3D GRE EPI (head gradient 100/1200, theoretical) |  | 0.575 | 0.575 | 0.575 | 220 | 220 | 46 | 382.6087 | 382.6087 | 80 | 40 | 0.07 | 2.8 | 4182.5547 | 100/1200 head-gradient system, using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total
3D GRE EPI (head gradient 200/900, theoretical) |  | 0.525 | 0.525 | 0.525 | 220 | 220 | 42 | 419.0476 | 419.0476 | 80 | 40 | 0.07 | 2.8 | 5017.1688 | 200/900 head-gradient system (Feinbergatron), using TE=T2* (30ms), 5ms for excitation, TAQ 50 ms, R=4, 10 ms for fat suppression, 5 ms for spoiling => 70 ms TR total